# PH405C pH Sensor — Arduino Firmware Documentation

## Overview

This firmware reads analog voltage from a PH-4502C pH sensor module connected to an Arduino UNO, converts it to a digital pH value using a linear calibration formula, and prints the result over the Serial Monitor.

| Property | Value |
| --- | --- |
| Microcontroller | Arduino UNO |
| Sensor | PH-4502C pH sensor module |
| Analog Pin | A0 |
| Serial Baud Rate | 9600 |
| Sampling Method | 15-sample rolling average |
| Update Interval | ~1.5 s (plus sampling time) |

## Pin Configuration

| Signal | Arduino Pin |
| --- | --- |
| pH Sensor Analog Output | A0 |

## Global Constants & Variables

| Name | Type | Value | Purpose |
| --- | --- | --- | --- |
| PH_Pin | const int | A0 | Analog input pin connected to the sensor |
| V_Ref | const float | 5 | ADC reference voltage (Arduino UNO operates at 5V logic) |
| ADC_res | const float | 1023 | 10-bit ADC resolution ($2^{10} - 1$) |
| sensor_slope | float | -5.56 | Slope of the voltage-to-pH linear calibration equation |
| sensor_offset | float | 20.89 | Offset (intercept) of the calibration equation |

Calibration source: values referenced from cimpleo.com — Arduino pH Meter using PH-4502C

## setup()

- Initializes UART serial communication at 9600 baud.

- Prints a firmware banner to the Serial Monitor for identification:

===================================

---------PH405C Firm ware----------

===================================

## loop()

### Step 1 — Sampling

- Takes 15 analog readings (total_samples = 15) from PH_Pin.

- Each reading is separated by a 10 ms delay to reduce noise/jitter.

- All readings are summed into raw_adc_sum.

### Step 2 — Averaging

float average_adc = raw_adc_sum / total_samples;

Produces a smoothed ADC reading, reducing the effect of electrical noise from a single sample.

### Step 3 — Voltage Conversion

float voltage = average_adc * (V_Ref / ADC_res);

Converts the averaged 10-bit ADC value (0–1023) into an actual voltage (0–5V) based on the Arduino's reference voltage.

### Step 4 — pH Calculation

float caluculated_ph = (sensor_slope * voltage) + sensor_offset;

Applies the linear calibration equation:

$$\text{pH} = (\text{slope} \times \text{voltage}) + \text{offset}$$

### Step 5 — Output

Prints the following to the Serial Monitor every cycle:

- Raw ADC average (1 decimal place)

- Sensor voltage (3 decimal places)

- Final calculated pH value (2 decimal places)

Then waits 1500 ms before repeating the loop.

## Practical Mistakes & Key Learnings

- Manufacturer Datasheet & Pinouts: Identify the exact manufacturer specifications and probe pinouts prior to usage, as variations exist across third-party suppliers.

- Warm-up / Stabilization Delay: Always allow at least a 60-second stabilization window every time the sensor is powered on or immersed in a new solution before taking measurements.

- Protective Storage Cap: Remember to remove the white protective cap (containing storage solution/gel) covering the glass probe tip before taking readings.

- Static Electricity Avoidance: Do not wipe or rub the glass bulb/sensor body with a dry cloth or tissue. Wiping accumulates static charges across the housing, leading to drifting or unstable ADC values. Rinse with distilled water instead.

- Shorting Connector Test: Short the BNC connector on the module board prior to initial setup to verify the $2.5\text{ V}$ mid-point reference baseline (which corresponds to neutral $\text{pH } 7$).

## Suggested Improvements

- Serial output formatting — current Serial.print() calls run together without line breaks or spacing, producing hard-to-read output like:

raw ADC Avg: 512.3sensor voltage: 2.500Final PH value7.06

Consider adding \t or Serial.println() between fields, e.g.:

Serial.print("Raw ADC Avg: "); Serial.print(average_adc, 1);

Serial.print(" | Sensor Voltage: "); Serial.print(voltage, 3);

Serial.print(" | pH: "); Serial.println(caluculated_ph, 2);

- Variable naming — caluculated_ph has a typo (should be calculated_ph); harmless but worth cleaning up for readability.

- Calibration verification — since sensor_slope/sensor_offset were sourced externally, recalibrate using pH 4.0, 7.0, and 10.0 buffer solutions for accuracy with your specific sensor unit.

- Temperature compensation — pH readings are temperature-sensitive; consider adding a temperature sensor (e.g., DS18B20) and compensation formula for higher accuracy if used outside controlled conditions.